## Slide 1
Reconstructing the 12-lead ECG
from a reduced set of leads
A masked-lead transformer trained on PTB-XL | 21,799 clinical ECGs | 4.9M parameters
Notes: The question: how few electrodes can you place and still recover a diagnostically useful 12-lead ECG? Trained one model to handle any subset of leads, then measured where it succeeds and where it fails.

## Slide 2
Approach
One model, any electrode configuration
01
Masked-lead training
Randomly hide leads each sample, 1 to 7 visible. The model learns every electrode configuration at once instead of one fixed set.
02
Predict 8, derive 4
III, aVR, aVL and aVF are exact linear functions of I and II. The network predicts only the 8 independent leads; the rest follow by algebra.
03
Score hidden leads only
Correlation averaged over all 12 includes leads handed to the model as input. Those score ~1.0 and hide the real reconstruction quality.
Einthoven and Goldberger identities hold exactly (error 0.0), because the dependent leads are computed rather than predicted.
Notes: Predicting all 12 wastes capacity reproducing arithmetic and yields traces that violate Einthoven's law slightly — visibly wrong to a cardiologist.

## Slide 3
Accuracy scales with electrode count
Held-out test set, fold 10 (2,198 records). Correlation on hidden leads only.
Chart: 
| 1 lead | 2 | 3 | 4 | 5 | 6 | 7 leads
Hidden-lead correlation | 0.813 | 0.874 | 0.901 | 0.915 | 0.924 | 0.928 | 0.934
Where the gains come from
1 → 2 leads: +0.061
The largest single jump. One extra electrode.
4 → 7 leads: +0.019
Three more electrodes buy a third as much.
Diminishing returns set in at 4 leads.
Notes: Reported number is Pearson r on leads the model could not see.

## Slide 4
Which leads, not how many
Exhaustive search over all 28 two-lead and 56 three-lead subsets
Two limb leads
I + II
r = 0.856
Both lie in the frontal plane — they span only two dimensions.
Two chest leads
V2 + V5
r = 0.910
Spread across the transverse plane — the missing third dimension.
The principle is spatial diversity, not lead count.
The worst subsets pair electrodes pointing the same way — I+V6, V5+V6, I+II. The best spread across septal, anterior and lateral positions.
Hardware caveat: precordial leads are referenced to the Wilson central terminal (average of RA, LA, LL), so V2 cannot be recorded without the three limb electrodes — which give I and II for free. Rank configurations by electrode count, not lead count.
Notes: Best buildable 4-electrode configuration is I+II+V3 at r=0.925, which beats I+II+V2. If you add one chest electrode, put it at V3.

## Slide 5
Clinical accuracy: ST-segment error
Correlation is dominated by the QRS complex and barely sees the ST segment — which is what indicates ischaemia.
Chart: 
| Normal | Myocardial infarction | ST/T change | Conduction disturbance | Hypertrophy
ST error (mV) | 0.0295 | 0.0475 | 0.0439 | 0.0523 | 0.0545
6.0%
of records exceed the 0.1 mV
threshold at 1 lead
0.5%
at 4 leads
(I+II+V2+V5)
Pathology costs accuracy
Abnormal records carry 1.5–1.9× the error of normal ones. PTB-XL is 44% normal ECGs, so any dataset-wide average understates error on exactly the cases that matter.
Notes: Measured at I+II. Conduction disturbance and hypertrophy are worst — both alter voltage relationships between leads in ways a model trained mostly on normal conduction has no basis to reproduce.

## Slide 6
Known problems
Found by inspecting reconstructions, not by reading metrics
Amplitude collapse
Predictions are visibly flattened on hidden precordial leads. L1 loss rewards hedging toward the conditional median. Correlation is scale-invariant and cannot detect this at all.
Lead II is the weakest lead
r = 0.82, below every precordial. Predicting it needs the frontal QRS axis, which varies −30° to +90° across patients and is unobservable from one lead.
Tail risk, not mean error
Single-lead ST error averages 41 μV but reaches 107 μV at p95 and 132 μV for infarction — above the 100 μV diagnostic threshold.
Reconstruction is indistinguishable from measurement
A synthesised trace looks identical to a recorded one. A clinician cannot tell which leads were inferred — a real risk of false confidence.
Notes: Amplitude collapse was spotted in a plot after several scalar metrics had all reported that things were fine.

## Slide 7
Where this goes next
Ranked by expected value, highest first
Patient-specific calibration
One simultaneous 12-lead recording per patient, used to fit a personalised output layer. Much of the residual error is anatomical — heart position and torso geometry — which is constant per person and learnable from one snapshot.
Highest expected gain
Amplitude-aware objective
Add a term penalising per-lead standard-deviation mismatch, the only part of the loss that punishes shrinkage. Implemented; awaiting an A/B run.
In progress
Downstream classification
Three-way test: classifier on true 12-lead, on reconstructed 12-lead, and on raw 2 leads. Determines whether reconstruction adds information or is an expensive identity function.
Decides the architecture
Pretrain, then fine-tune
Masked-lead prediction is a strong self-supervised objective. Keep the encoder, discard the decoder, fine-tune for diagnosis directly.
Strongest variant
Notes: Scaling to MIMIC-IV-ECG (~800k records) for pretraining is the other large available win, and 500 Hz data would test whether the 100 Hz result holds.

## Slide 8
What holds up
Rhythm
Usable from 1 lead
Rate, atrial fibrillation, ectopy and conduction intervals need one clean lead. Already an FDA-cleared consumer capability.
Morphology
Needs 3–4 leads
r = 0.925 at I+II+V3, four electrodes. Comparable to EASI-derived 12-leads used in bedside monitors.
Ischaemia
Needs 4+ leads
ST error falls below a fifth of the diagnostic threshold only at I+II+V2+V5. At 1–2 leads the tail is unsafe.
5 electrodes reconstruct a diagnostically usable 12-lead ECG. 2 electrodes do not.
PTB-XL fold 10, n = 2,198 | 100 Hz | transformer, 4.9M parameters | trained on a single RTX 4050
Notes: Ten-electrode carts remain the standard; this halves that count.